ESA614 / ESA615 Update Instructions

Firmware version 3.01

Instructions for updating your ESA614 or ESA615:

Note: The instructions are the same except for the SD card image used in Step 8.

- If necessary, download and install Ansur Executive 2.9.7 (or later) and Ansur ESA615 Plug-In 1.0.5.1 (or later) from http://www.flukebiomedical.com/Biomedical/usen/Support/software/Ansur-Software.htm

- Extract the files from this zip archive.

- (Mandatory) Update SD card with new factory image by following Steps 4 through 15.

- Insert SD card into an SD card reader (not included) and plug into your PC.

- Back up any custom test sequences (files ending in “.tem”) and all test results (files ending in “.dta”) to your PC.

- Run “snapshot.exe” program which was extracted from the zip archive.

- Click “Restore Disk from File” button.

- Click “Browse” and find the file “ESA615_SD_Image_FW_3_01_20180723.SNA” or “ESA614_SD_Image_FW_v3_00_20180723.SNA” extracted from the zip archive.

- Click “Next”

- Select the SD card’s drive letter. Ensure the Size is 484 or 470.

- Click “Next”

- Click “Yes”

- Click “OK”. The error does not need further action.

- Click “OK”

- Click “Exit”

- (Optional) Convert test results, as desired:

  - Run the ESA615 Mini Plug-In

  - Click the “…” icon

  - Select the format to convert to

  - Click Convert

- Run the Update Firmware tool via the Start Menu > All Programs > Fluke > Update Firmware.

- Select ESA615 (ESA615 is also used for ESA614) and click Next.

- Click the first “…” button associated with UI Firmware file and find “ESA615_UI_v3_01.hex” which was extracted from the zip archive.

- Click the second “…” button associated with DSP Firmware file and find “ESA615_MTR_v3_01.hex” which was extracted from the zip archive.

- Click the third “…” button associated with VNC Firmware file and find “ESA615_VNC_v1_09.rom” which was extracted from the zip archive.

- Click Next.

- Plug the ESA614 or ESA615 into your PC and power it on.

- Click Upload.

- After the upload finishes, your unit should boot normally. No calibration is required if upgrading from at least firmware v2.12. Older firmware requires a recalibration!